## Slide 1: Development & Validation of BNL’s TBD for Release of M&E from Accelerator Facilities Chuck Schaefer, CHP Accelerator Safety Workshop September 2019

## Slide 2
NSLS2 has ~ 15 Credited Controls
National Synchrotron Light Source II Facility

## Slide 3
CAD has ~ 40 Credited Controls
Collider – Accelerator Complex

## Slide 4
In early 2017 RCD Management decided to pursue TBD
Focus only on volume activity
Tier 1 (IFB only) to be used for routine operational releases
RCD’s existing survey process for M&E was to use Micro-R Meters
99% of BNL’s volume-activated material comes from C-AD

## Slide 5: Survey Process
On-Contact scan surveys using Micro-R Meters in low-background areas < 13 µR/h
Scan at ≤ 2 inches/sec
Cover all accessible surfaces
No formal direction provided for surveying “small” items

## Slide 6
PROBLEM #1: Where did 13 µR/h come from?
Had been limiting BKG value for performing activation surveys since 1990s
NO documentation existed for its basis:
 Lessons Learned #1
Recent analysis showed can’t reliably detect volume activity of Group one R-Nucs (3 pCi/g) in a BKG of 13 µR/h

## Slide 7: PROBLEM #2: What is an Acceptable Value of BKG?
Analyzed MDCR corrected for surveyor efficiency & converted to µR/h
6 µR/h is limiting BKG for reliably detecting 3 pCi/g of proxy R-Nucs (60Co, 22Na, etc.)
MDCs calculated using Microshield for objects from 0.01 Kg - 100 Kg

## Slide 8
Chart: 60Co Activity Concentration for 2 µR/h Detection Threshold
| 10 | 100 | 1000 | 10000 | 100000
Activity Concentration | 37 | 8.7 | 2.4 | 1.1 | 0.8
Screening Level | 3 | 3 | 3 | 3 | 3

## Slide 9: PROBLEM #3: Micro-R Meter a Good Release Instrument?
Spent 2 months performing a site-wide BKG survey & instrument comparison
Used 3 Scintillators: Micro-R Meter, Bicron Micro-Rem, & Radeye
Micro-R Meter exhibited best response behavior; lowest CV: Selected for use
Many areas ≤ 6 µR/h; Operations not restricted

## Slide 10: Back to the TBD
Asked 3 authors of DOE-STD-6004 to peer review the TBD
Great comments!...but painful also
Had to develop a basis for maximum 8-inch distance between survey paths for large items  Instrument Sensitivity Study
Had to develop facility-specific Bounding Conditions (e.g., how long is 60Co an adequate proxy for HTDs?)

## Slide 11: Approval Process
Addressed all reviewer comments and internally approved in November 2017
DOE-BHSO forwarded to HQ for their review (J. Blaikie and M. Epps)
Only a couple questions, but no revision necessary
DOE-BHSO Formally approved in January 2018

## Slide 12: Post-Approval Validation
Six RCTs individually surveyed 8 low-activity, small metallic items (0.3 – 3 kg)
Observations recorded as either “Rad” or “Non-Rad”
Samples subsequently characterized by gamma spectroscopy and offsite analysis for HTDs
Probability of detection plotted using dose response model

## Slide 13
Chart: Probability of Detection (P) for Low-Mass Items
| 1.2E-2 | 3.1079999999999997E-2 | 7.3999999999999996E-2 | 0.19239999999999999 | 0.49209999999999998 | 0.59570000000000001 | 0.24419999999999997 | 0.29599999999999999
 | 0 | 0 | 0.66667 | 0.83333 | 1 | 1 | 0.66667 | 1

## Slide 14: Summary
Instrument and site BKG studies proved very helpful
Establish Max. BKG for performing release surveys early in TBD development
Consider grouping small-mass items for survey whenever “practicable” to counter the mass effect
Peer review is a must!!!
Involve RCTs in the process

## Slide 15
Thank you
Questions or Comments?