## Slide 1: Multi-Grid Boron-10 detector: from proof-of-principle to applications in time-of-flight spectrometers
Anton Khaplanov
European Spallation Source ERIC
15/10-2019

## Slide 2: Outline
Multi-Grid technology
  Proof-of-principle
  Characterization
Prototypes and technology demonstrators
  IN6
  CNCS
  SEQUOIA
Results
  Results from chopper spectrometer tests
  Recent measurement on V20

## Slide 3: Acknowledgements
Multi-Grid proposed by the ILL. Developed collaboratively with ESS and Linköping University
Tests at ILL, SNS, HZB, IFE, LLB
The bulk of development under

## Slide 4: Development Requirements
Alternative to the existing 3He technology
Original driver – cost
Further driver – rate capability
Detector performance needs:
  Efficiency
  Background
  Gamma rejection
  Fast neutron sensitivity
  ToF/energy resolution
  Uniformity
Design needs:
  Cost
  High active area
  Vacuum compatibility

## Slide 5: Multi-Grid Technology
Detection in multiple thin layers of 10B4C
B-10 film
Al substrate
n
conversion
gas
fragment 1
fragment 2
gas
n
n
Low cost for large area
High total efficiency
Readout split over many layers – lower local rate
No need for resistive readout – low gain

## Slide 6: Efficiency Optimization
Many B-10 layers that don’t have to be the same
1.5μm
0.5μm
1.0μm
n
Efficiency depth profile,
efficiency per blade
 Constant vs. optimized B4C thickness
Grid with varied B4C layer thickness
Efficiency optimized depending on wavelength prioritization

## Slide 7: B4C Layer Thickness
Choices for cold and thermal instruments
1.5μm
0.5μm
1.0μm
n
7 blades 0.5μm,
7 blades 1.0μm,
3 blades 1.5μm
4 blades 1.0μm,
10 blades 1.25μm,
6 blades 2.0μm.
Efficiency as a function of wavelength for MG optimized for cold vs. thermal spectrum
Layer thickness in MG.SEQ (20 blades)
 | CSPEC | T-REX
Typical initial λ, Å (meV) | 2 to 15 Å (20 to 0.36 meV) | 0.7 to 6.4 Å (160 to 2 meV)
CSPEC-optimized grid
T-REX optimization
MG.SEQ
CSPEC optimization
MG.CNCS
Layer thickness in MG.CNCS (16 blades)

## Slide 8: 10B4C Layer Production
Magnetron sputtering in ESS workshop in Linköping
DC magnetron sputtering used to create coatings on Al and other substrates
97% enriched 10B4C
Layer thickness 0 to about 10um
Composition, impurities:
Produced blades for MG
Thickness uniformity
Deposition on blades

## Slide 9: Prototype 1
Built and tested at ILL – demonstration of detection and efficiency
Pulse height spectra due to conversion process with 2 particles
Measured confirmation of efficiency
function of:
Number of layers (here = 15),
Thickness of layers,
Threshold,
Neutron wavelength (here = 2.5Å)

## Slide 10: Prototype 2
Built in ILL, demonstrated operation of long detector
2-m stack of 96 grids with 64 wires
Technological feasibility
Deformation effects
Beam test at CT2 in ILL
A
B
C
D
Uniformity study

## Slide 11: IN5 Demonstrator
Demonstration of full-scale detector production
Detector:
2.4 m2 active area
Blade x18432:
enriched B4C coating
good adhesion, uniformity,
Column x8:
wire-frame coincident readout
Grid x1024:
low background,
minimal dead material
80cm
3 m
IN5 Spectrometer at ILL

## Slide 12: IN6 Demonstrator
First demonstration on ToF spectrometer
96 grids in 6 modules of 16
30x45 cm2 active area
Installed and operated at IN6 at ILL for 2 weeks
Data compared to 3He detectors at the same angle

## Slide 13: IN6 Demonstrator
An unexpected issue – intrinsic background
25 um Ni plating or pure Al
 x100 improvement
25 um Ni plating
pure Al (5N5+ purity):
U/Th concentration < 1ppbw

## Slide 14: MG.CNCS Installed at CNCS, SNS
MG
He-tubes
1.1m
A.Khaplanov et al. “Multi-Grid Detector for Neutron Spectroscopy: Results Obtained on Time-of-Flight Spectrometer CNCS" https://arxiv.org/abs/1703.03626 2017 JINST 12 P04030
Long-term functionality of MG on a cold spectrometer

## Slide 15: MG.CNCS Demonstrator
Design and assembly
15
1.1m
96 grids in 2 modules of 48
110x19 cm2 active area
3-thickness optimized grids in construction
Installation of anode wires
64 wires per module
20-um tungsten Au-plated wire

## Slide 16: MG.CNCS Results Summary
Snapshot from 1 year of operation
Vanadium
Background
Single crystal
Fast neutrons (fission, prompt)
Water, Ice
Allowing to measure:
Energy resolution, efficiency, scattering, background sensitivity, saturation
Gamma background
(NaI detector installed near MG)

## Slide 17: MG.SEQ Detector Tested at SNS
Test with thermal n; first test in vacuum
Characteristic | Values
Detector area | 0.8 m2
# of grids | 360*
# of wires | 720
# readout channels | 1080*
# voxels | 28800
# of 10B4C layers | 40
10B4C on radial blades | Yes + No
# of detectors | 3 (+1 spare)
Detector weight | 110kg x3
Sample-detector | ~5.5m
Detector angle | 25-35 deg
Electronics compartment
Simplified module access

## Slide 18: MG.SEQ Testing
In workshop, in beam, vacuum/pressure
Vacuum / pressure tested at EWCON
Beam test at ILL, CT2
Neutron Beam/ source tests for detection
Tested to 1.1 bar overpressure,
10+ cycles
Leak rate measurements at 1 bar overpressure
Vacuum tank test at SNS

## Slide 19: MG.SEQ at SEQUOIA
MG test in vacuum and at high energies
White beam on Si-powder
Vacuum connections to the outside
The three detectors installed in the lower detector bank
T0
Sample
Fermi
Monitor
MG.SEQ
Primary spectrometer provides beams with Ei from 2 to 5000 meV

## Slide 20: Data Acquisition and Analysis
Spectra and their features
Time-of-Flight
Energy Transfer

## Slide 21: MG.SEQ Measurements: Vanadium
Comparison to He-3 at a range of energies
50 meV
8 meV
25 meV
160 meV
80 meV

## Slide 22: MG.SEQ Measurements: C4H2I2S
Sample with vibrational and rotational levels over a range of E
50 meV
100 meV
70 meV
200 meV
Spectra at several Ei
MG vs. 3He
Agreement in peak position and width over the range of Ei and Ef

## Slide 23: MG.CNCS/SEQ Elastic Enerrgy Resolution
FWHM of the elastic line
MG.CNCS
MG.SEQ
Good agreement in FWHM for all energies in HR and HF modes
MG slightly higher resolution – likely due to thinner detector cells

## Slide 24: MG Efficiency as a Function of Energy
Measured relative the 3He detectors at CNCS and SEQUOIA
Counts from integral of the elastic peak in dE.
Efficiency measured relative to He3 tubes at the same scattering angle.
He3 efficiency according to the calibration curve used in the SNS instrument used for normalization.
MG.CNCS
MG.SEQ

## Slide 25: Multi-Grid Gamma Sensitivity
… is very low, similar to He-3 detectors
Gamma sensitivity in the IN6 measurement
Measured with 164 MBq 137Cs source, γ (662 keV)
Multi-Grid He-3
Effect of the threshold on neutron efficiency
* A. Khaplanov et al., JINST 8, P10025 (2013), arXiv:1306.6247

## Slide 26: Rate Study in MG.CNCS
Using a Bragg Reflection to create an intense localized spot
Crystal sample
Incident beam
Reflected beam at Ei = 17.20meV
Reflected beam at Ei = 13.74meV
He3
MG
He3
MG
Low intensity
High intensity
Loss of position resolution
Position:
Charge division in He3 PSD – saturation causes loss of position resolution
Coincidence between grids and wires in Multi-Grid
Local count rate limit:
space charge in gas (*)
  x10 lower in MG due to multiple layers
  x10 lower in MG due to lower gain
Dead time of electronics
depends on electronics, similar

## Slide 27: MG Background
Shielding for low-level investigations in Utgård
Shielding cave:
Mirrobor, 5mm, (80% B4C)
HDPE, 70mm
Detector | Conditions | Background rate, Hz/m2
3He CNCS | Source off, shutter closed | 0.4
3He SEQUOIA | Source on, shutter closed | 0.15
MG.CNCS | At CNCS, source off shutter closed | 0.55 (no cuts) 0.25 (with cuts)
MG.SEQ (pure Al) | At CNCS, source on shutter closed | 1.40 (no cuts) 0.63 (with cuts)
MG.SEQ (pure Al) | In Utgård | 0.3 (no shielding) 0.07 (with shielding) 0.04 (with shielding and cuts)
Background measurements available in many configurations.
Large differences due to location and shielding
Some regions had background issues

## Slide 28: Scattering in Grids
Why a ToF measurement is necessary
Looking for a <10-2 effect
Continuous beam measurements are not sufficient. Problems arise from:
Beam halo
Scattering along the direction of the beam
Incomplete shielding
Background
IN6 Data
Detector depth
ToF

## Slide 29: Complications with Scattering Measurements
Only accessible in ToF, but many effects complicate results
Incident n
Sample
Sample environment
Detector
Scattering effects from SE and detector
The shoulder is not always easy to quantify consistently.
Many energy-dependent effects influence the shape simultaneously. Many of them – not detector-dependent.
When scattering “switches on” in MG, it also switches on in other beam-line components.

## Slide 30: MG.SEQ Results: Scattering Effect on Line Shape
Line shape comparison to He3
 Scattered neutrons appear slower than they should.
Effect of scattering is energy dependent.
Appears on energy-loss side in dE spectra.
Define as relative intensity of the shoulder between 3𝜎 and 5𝜎.
Effect of scattering as a function of energy for the 3 MG.SEQ modules and the SEQUOIA He3 array.
Trend – coating the radial blades reduces the effect.

## Slide 31: HZB V20 Beam Line Measurement
Direct beam monochromated with a Fermi chopper
Direct pulsed beam measurement
Fermi chopper to generate a train of monochromatic pulses
Pulse resolution ~10us
Interval 1 ms
Study after-pulse effects, broadening due to scattering in grid
Monitor
Detector
Fermi chopper
Collimators

## Slide 32: Further In-Beam Tests
V20 direct beam with Fermi chopper
Direct pulsed beam measurement
Fermi chopper to generate a train of monochromatic pulses
Pulse resolution ~10us
Interval 1 ms
Study after-pulse effects, broadening due to scattering in grid
*note: shielding around detector was added after this photo

## Slide 33: Conclusion
Development Requirements addressed
Alternative to existing 3He technology  Detector validated for ESS ToF Spectrometers
Original driver – cost  Cheaper than 3He
Further driver – rate capability  Exceeds 3He
Detector performance needs:
  Efficiency
  Background
  Gamma rejection
  Fast neutron sensitivity
  ToF/energy resolution  Matching, or close to 3He
  Uniformity
Design needs:
  Cost
  High active area
  Vacuum compatibility
Studies of the detailed line shape are ongoing