## Slide 1: Microscopic samples and high precisions: challenges and solutions at the ESS
Malcolm Guthrie
European Spallation Source ERIC

## Slide 2: The Challenge
Key design feature of recent neutron sources is beam brightness
Allows significant reduction in sample sizes.
Requires step change in precision and repeatability of sample positioning
Chart: Minimum Envisaged Sample Size @ ESS
| <1MM | <1CM | >1CM
Minimum Envisaged Sample Size | 8 | 6 | 1
Sample diameter ~ 500μm

## Slide 3: Kinematic mounting
ESS has defined a standard mounting system based on kinematic positioning
Geometric constraints systematically constrain 6 degrees of freedom
Reproducibility < 1 μm (theoretical)

## Slide 4: Universal Coordinate System
Absolute repeatability enables concept of “universal coordinate system”:
  define origin as sample position
  define cartesian axes with standard convention, aligned to beam
  map this system onto all beamlines
  Can also map onto offline stations (alignment, Raman...)

## Slide 5: Alignment protocol* (*assumes known sample position “+”)
Inst. 1
Inst. 2
Inst. “n”
Before alignment:
Beam
Calib’Standard

## Slide 6: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Beam
Calib’Standard

## Slide 7: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Beam
Calib’Standard

## Slide 8: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Beam
Calib’Standard

## Slide 9: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Beam
Calib’Standard

## Slide 10: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Etc…
Beam
Calib’Standard

## Slide 11: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Etc…
Beam
Calib’Standard

## Slide 12: Alignment protocol* (*assumes known sample position!)
Inst. 1
Inst. 2
Inst. “n”
Install standard
Adjust mount to align
Repeat on Inst 2
Etc…
Beam
Calib’Standard
Expect periodic recalibration (beam drift, concrete subsidence etc.)
Note: Samples must be calibrated relative to “equipment” mount
This can be done offline by calibrating an “alignment station” in exactly the same way as the instruments

## Slide 13: Universal Coordinate System
Allows rapid exchange of different equipment on one beamline
Allows rapid exchange of same equipment on many beamlines
Allows offline definition of ROI’s that can be transferred to beamline
Inst. 1
Inst. 2
Inst. “n”
Beam
Calib’Standard
Calib’Standard
Calib’Standard

## Slide 14: Kinematic component
Tested two alternative kinematic components:
Kipp “Zero point” system
Baltec Canoe Spheres/Vee blocks

## Slide 15: Kinematic component
Tested two alternative kinematic components:
Comparison
Mechanism | Pros | Cons
Baltec | Available with multiple material choices all interface components are identical relatively low cost for one-off instrument part of interface (€450) | Relatively high ongoing cost (will cost €450 for each item of SEE) Small scale "family business" manufacturer based in CA, USA, medium risk to long term supply of standard. Only provides alignment, requires separate clamping mechanism standard material is magnetic, non-magnetic is available but at additional cost
Kipp | Well-established supplier, low risk of long-term supply of standard. built-in clamping mechanism Relatively low ongoing cost (will cost €90 for each item of SEE) | magnetic component and mfg hasn't yet confirmed if non-magnetic version is possible (although we're in discussion) Relatively high cost for one-off instrument part of interface (€1200).

## Slide 16: Test setup
Mounted both “instrument” interfaces, rotated by 30°
Mounted matching “equipment” interfaces onto flange

## Slide 17: Test setup
Installed 3 laser fiducials on top of flange

## Slide 18: Test setup
Both systems showed repeatability within tolerance of measurement
Chart: Installation Repeat
| 0 | 1 | 2 | 3 | 4
Point1 | 0 | 0.011 | 0.013 | 0.015 | 0.006
Point2 | 0 | 0 | 0.004 | 0.01 | 0.01 | 0.008
Point3 | 0 | 0.013 | 0.014 | 0.016 | 0.008
Will test effect of loading
  axial and offset
Effect of locking mechanism
Long term repeatability/wear

## Slide 19: Mapping the UCS onto sample position
Need to rapidly identify beam centre...
Use beam scanning methodology
y-axis
Beam
Detector
z-axis
Setup:
Position (mm)
counts

## Slide 20: HZB alignment, scanning tests
Set up
HZB beam monitor
X-Y-Z stage on SAD standard kinematic mount
Upstream slits (10x10mm)
Downstream slits (6x6mm)
ø25mm hBN disc (several pinhole sizes)

## Slide 21: HZB alignment, scanning tests
Position (mm)
counts
Interfaces
stages+
pinhole
Beckhoff driver
EPICS control layer
HZB monitor
Monitor readout
NICOS control layer
SAD
Motion
Control +ICS
V20 team
DMSC /
Det. Group
DMSC
Timing system
ECDC
Notes: MESI had critical role.
Encountered sharply defined and well-defended scope (e.g. motor control). similar to instrument experience.

## Slide 22: Multiple installation levels

## Slide 23: Sister system: “Flange mount”
Design for Level 1 interface finished
Prototype support frame ready for procurement
Detailed design for Levels 2-3 ongoing
Level | Distance to beam (mm),Hn | Flange-to-beam (“bucket” design) (mm), Hn | Weight max. (kg) | Opening diameter (mm) | Proposed absolute ± positioning tolerance at sample position (mm)
1 (XL) | 600 | 600 | 1000* | , | 1-2 TBC
2 standard (L) | 600 | 350 (TBC) | 300 | Φ505 | 0.25 TBC
3 high-precision | 600 | 200 (TBC) | 5 | Φ255 | 0.01 TBC
Level | Distance to beam (mm),Hn | Flange-to-beam (“bucket” design) (mm), Hn | Weight max. (kg) | Opening diameter (mm) | Proposed absolute ± positioning tolerance at sample position (mm)
1 (XL) | 600 | 600 | 1000* |  | 1-2 TBC
2 standard (L) | 600 | 350 (TBC) | 300 | Φ505 | 0.25 TBC
3 high-precision | 600 | 200 (TBC) | 5 | Φ255 | 0.01 TBC

## Slide 24: Acknowledgements
Lauritz Saxtrup
Anders Pettersson