## Slide 1: Development of a Polarized 3He++ Ion Source for the EIC
Matthew Musgrave
G. Atoian, E. Beebe, S. Ikeda, S. Kondrashev, J. Maxwell, R. Milner,
M. Okamura, A. Poblaguev, D. Raparia, J. Ritter, S. Trabocchi, A. Zelenski
PSTP 2019
September 26, 2019

## Slide 2: Why a Polarized 3He Ion Source?

## Slide 3: Source Concept
Identified as high priority R&D for EIC by EICAC review in 2009, Office of Nuclear Physics Community Review in 2017, and the 2018 assessment of the US National Academy of Sciences.
Requirements
Polarized 3He using optical pumping and injection into EBIS at 5 T
Maximum polarization >70%
Intensity 2.5 × 1011 3He++ ions in 20 𝜇s pulse (~4 mA peak current)
Spin-flip in the beam transport line

## Slide 4: RHIC EBIS
5T solenoid B Field; 1.5 m ion trap
20 keV electrons up to 10 A
5 Hz maximum repetition rate
5 T
Radial trapping of ions by the space charge of the electron beam
Axial trapping by applied electrostatic potentials at ends of trap
Ion output per pulse is proportional to the trap length and electron current

## Slide 5: Extended EBIS Upgrade
Electron Gun
Electron Collector
Add a second 5 T superconducting solenoid for trap length extension: 40% increase in Au intensity
Install 3He polarization and injection system into the upstream solenoid
Opportunity for various vacuum and other minor improvements
3He polarization and injection systems

## Slide 6: Extended EBIS Superconducting Solenoids
5.0 Tesla field, about 1.0 Tesla at field minimum in solenoid separation (~30 cm)
Electron beam successfully propagated through both solenoids in May 2019.

## Slide 7: Polarized 3He Ion Source

## Slide 8: Metastability Exchange Optical Pumping

## Slide 9: High-Field MEOP
Lumibird Pump Laser Specifications
276769.46 GHz central frequency
2 GHz linewidth
~300 GHz thermal tuning range
5 GHz remote fine tuning range
5 W Power
GHz Shift from 3He C1 Frequency at 0 T

## Slide 10: RF Discharge at Multi-Tesla Fields
RF discharge parameters strongly affect maximum polarization.
RF discharge power needs to be reduced as 3He polarization increases.
Optimization of the 3He cell geometry and placement of RF electrodes should improve polarization.

## Slide 11: Optical Probe Polarimetry
GHz Shift from 3He C1 Frequency at 0 T
Pumping laser
Probe laser, Toptica, diode laser
Detector
Wavemeter
Pumping laser-Keopsys 10 W 1083 nm- fiber laser
3He cell

## Slide 12: Open 3He Cell and Gas Purification System
Non-magnetic brass pneumatic remotely controlled Isolation Valve
OPPIS (RHIC’s polarized proton source) was converted into a high field MEOP system for polarizing 3He.

## Slide 13: Open 3He Cell and Gas Purification System
3He cyro-purification and storage system was built from a modified cryo-pump. Pumps everything except helium!

## Slide 14: 3He Open Cell Results
3He pressure can be controlled with the cryo-pump temperature.
Polarizations have been measured >80% at 2 torr and higher pressures have similar results.
Relaxation time of 30 s measured in open cell. Relaxation rate is limited by metal surfaces of the fill valve and gas contamination.
Valve construction and open cell design will be optimized to improve relaxation rate and maximum polarization.

## Slide 15: High Speed Pulsed 3He Valve
Pulsed current causes valve to open by Lorentz Force to the conducting plate in high magnetic field.
Valve design has successful long-term operation in OPPIS: B=3 T, I=100 A, L=5 cm, F=15 N, τ=100µs.

## Slide 16: High Speed Pulsed 3He Valve
Test of high speed pulsed valve in a 2.5 T field in the OPPIS solenoid.
B = 2.5 T, I = 12 A, 3He reservoir pressure = 2 Torr
µs
*Actual valve open time will differ from the electric current pulse width.

## Slide 17: Gas Ionization Cell for Electron Beam Tests in the BNL TestEBIS
46 cm long, 1 cm diameter gas ionization cell
3 cm long, 0.5 cm diameter down stream barrier (blue)
Electron beams of 6 Amps were successfully propagated through the assembly.

## Slide 18: Vacuum Simulations of EBIS with MolFlow
Electron Beam
1 mm diameter
10 Amp
25 keV
Timing & Trap Capacity
Maximum EBIS pulse time = 200 ms
Charge Capacity = ~1012 elementary charges
3He++ Capacity = ~ 2.5x1011 3He++
Pressure
1x10-6 mbar during gas injection
1x10-10 mbar standard operation

## Slide 19: Electron Beam Ionization of 3He
The e-beam acts like an ion pump in the EBIS vacuum; therefore, treat the e-beam as an ideal pump with a transparency defined the pumping efficiency (S).
The pumping efficiency is the ratio of the 3He mean free path through the e-beam and the average distance the 3He travels through the e-beam (from the electrons reference frame).
The actual pumping rate of the e-beam is independent of the beam radium & gas velocity because these are canceled out by parameters in the simulation.
3π/8 correction to average velocity of gas that hits e-beam.

## Slide 20: Electron Beam Ionization of 3He
For an e-beam of 25 keV there is a ≈0.5% probability that 3He is ionized during traverse of the e-beam.
Treat the e-beam as an ideal pump with 99.5% transparency.

## Slide 21: Electron Beam Ionization of 3He
For an e-beam of 25 keV there is a ≈0.5% probability that 3He is ionized during traverse of the e-beam.
Treat the e-beam as an ideal pump with 99.5% transparency.
Proportion of 3He ionized after 20 ms

## Slide 22: Electron Beam Ionization of 3He

## Slide 23: Gas Diffusion and EBIS Vacuum
2.65 x 1012 3He atoms injected
10 A, 25 keV e-beam
10-10 mbar EBIS vacuum

## Slide 24: 3He Wall Collisions Before Ionization

## Slide 25: Simulation Results for 3He Injection into EBIS
All results are encouraging for the project!
Data for gas injection will be collected after installation of the Extended EBIS in the summer of 2020.

## Slide 26: 6 MeV 3He++ Polarimeter
“Precision absolute polarimeter development for the 3He++ ion beam at 5.0-6.0 MeV energy”
Grigor Atoian
Friday at 13:40

## Slide 27: Summary
Extended EBIS upgrade is ongoing and electrons were propagated through two 5 T solenoids.
3He was successfully polarized to >80% in a 3 T field.
An optical probe polarimeter to measure 3He polarization was developed.
A prototype high-speed pulsed valve has been tested.
Gas injection & ionization simulations show promising results.
The 3He spin-rotator is being designed and equipment purchased.
The 6 MeV 3He polarimeter has been designed and the concept was tested with an alpha source.
Now all the parts need to be put together. We plan for partial installation in the summer of 2020 and complete installation followed by polarization measurements in the summer of 2021.

## Slide 28: Thank you for your attention
MIT-BNL Polarized 3He Ion Source Collaboration
G. Atoian, E. Beebe, S. Ikeda, S. Kondrashev, J. Maxwell,
R. Milner, M. Musgrave, M. Okamura, A. Poblaguev,
D. Raparia, J. Ritter, S. Trabocchi, A. Zelenski
This work was supported by DOE Office of Nuclear Physics, R&D for Next Generation Nuclear Physics Accelerator Facilities.