## QIKR Motion WBS Components
FMEA
Process/Product Name: |  | QIKR Motion WBS Components |  |  |  | Quality Level: |  |  | 2
Responsible: |  | Danielle Wilson |  |  |  | FMEA Date (Orig.): |  |  | 2024-04-26 | (Rev.): | 0
[6 empty] | Prepared By: |  |  | Danielle Wilson
Process Step/Input/System | Potential Failure Mode | Potential Failure Effects | SEVERITY (1 - 10) | Potential Causes | OCCURRENCE (1 - 10) | Current Controls | DETECTION (1 - 10) | RPN | Potential Action Recommended | Resp. | Actions Taken | SEVERITY (1 - 10) | OCCURRENCE (1 - 10) | DETECTION (1 - 10) | RPN
What is the process step, change or feature under investigation? | In what ways could the step, change or feature go wrong? | What is the impact on the customer if this failure is not prevented or corrected? |  | What causes the step, change or feature to go wrong? (how could it occur?) |  | What controls exist that either prevent or detect the failure? |  |  | What are the recommended actions for reducing the occurrence of the cause or improving detection? | Who is responsible for making sure the actions are completed? | What actions were completed (and when) with respect to the RPN?
Incident Table | Need more vertical adjustment in the incident table than is currently available | Cannot select correct beam path | 10 | The beam paths do not split at the location expected | 1 | 1) Slits and frame overlap mirror can be manually adjusted vertically to some degree (10mm) 2) Incident table can be manually adjusted vertically to some degree (50mm) 3) Detector can tell if the beam is not selected correctly | 1 | 10 | None | [5 empty] | 0
Incident Table | Lead screw buckles under load | Incident table cannot be used until repaired | 10 | User puts heavy equipment on incident table, leans on table | 5 | Lead screw is designed to handle 200 additional lbs of loading beyond nominal with a safety factor of >3x | 1 | 50 | Could have a spare leadscrew on hand (for operations, not initial installation) | [5 empty] | 0
Incident Table | Incident table cannot be installed with the guide present | Delays start of beamline operation till problem is corrected | 10 | Not enough clearance between the guide and incident table for assembly | 2 | 1) Critical dimensions will be chekced before assembly 2) Monorail crane will enable easy component placement 3) 6mm of clearance will be allowed between the guide and the table component that fits around it | 1 | 20 | None | [5 empty] | 0
Incident Table | Lead screw binds | Cannot select correct beam path | 10 | Overconstraint, foreign material on screw, lack of lubrication (maintenance) | 1 | 1) Current pivot motion is constrained by the pivot bearings, not the lead screw 2) Lead screw is easily accessible for inspection & maintenance | 1 | 10 | 1) Specify maintenance schedule for checking/reapplying lubrication 2) Shield the leadscrews from accidental contact with foreign material | Rudy Thermer |  |  |  |  | 0
Incident Table | Pivot bearing binds | Cannot select correct beam path | 10 | Two halves of bearings are not well-aligned, lubrication issue (maintenance), dirt/fluids get into the bearing | 3 | 1) The two halves of the bearing housing will be machined in one setup to ensure coaxiality, dimensions will be measured and confirmed 2) one side of the bearing can float in the axial direction 3) There is a mechanism to slide the table out for better access to the pivot bearing to allow routine maintenance activities | 2 | 60 | 1) Specify maintenance schedule for checking/reapplying lubrication 2) Put a dust shield around the bearings | Rudy Thermer |  |  |  |  | 0
Incident Table | Foreign object is placed in the beam near the guide exit | Creates an unsafe radiation environment in the neighboring cave (QIKR-U or QIKR-L) | 8 | The user might accidentally leave an object (a tool, etc) near the end of the guide (though this would be difficult to do) | 2 | padlocked cover exists over the guide end of the incident table to prevent accidental placement of foreign objects | 10 | 160 | Install radiation monitors to detect spikes | [5 empty] | 0
Incident Table | Third slit pnuematic lock fails | user cannot move the third slit | 8 | pneumatic air supply fails, connection of air supply fails | 3 | May be able to overcome the lock with a tool (key) | 5 | 120 | Check that the key will be effective till repairs are effected | Rudy Thermer |  |  |  |  | 0
Incident Table | Sample environment could hit the third slit | slit or sample environment could be damaged | 10 | Slit is too close to the sample environment when remote sample setup begins, user angles/positions the SE such that it hits the slit | 5 | None currently | 10 | 500 | Compliant mount for the 3rd slit? Sensors that detect collision? Force feedback from the hexapod? | [5 empty] | 0
Incident Table | Motor fails | Axis cannot actuate, experiment cannot be set up | 10 | wire failure, connection failure, motor could overheat, lead screw binds | 2 | Components are conservatively designed with large safety margins | 2 | 40 | Could have spare motors on hand (for operations, not initial installation)
Incident Table | Encoder fails | Axis can actuate, but position cannot be verified within the motion software | 5 | Radiation damage, physical damage | 5 | 1) Using reflective optical encoders, which are less sensitve to radiation damage than thru-beam varieties 2) Placing the encoder tapes and read heads in locations where accidental damage is unlikely 3) The instrument detector is used to confirm correct placement of components via relative motions during experiment set-up procedure... absoulte position sensing saves time but is not necessary to complete sample setup | 1 | 25 | Limit the variety of encoder readheads and tapes, have spares on hand (for operations, not for initial installation)
Incident Table | Limit switch fails | Axis can actuate, but may overtravel | 3 | Out-of-the-box failure of the switch itself, physical damage | 3 | 1) Hard stops will be present and placed in locations to prevent major problems from overtravel 2) Absolute encoders will be present | 3 | 27 | Can use robust industrial type switches & test their operation prior to installation. Can have spares on hand (for operations, not for initial installation)
Sample/Detector Table | Lead screw for changing detector arm angle buckles under load | Table cannot be used until repaired | 10 | User puts heavy equipment on detector table, leans on table | 5 | Lead screw is designed to handle 200 additional lbs of loading beyond nominal with a safety factor of >3x | 1 | 50 | Could have a spare on hand (this would be for operations, not for initial installation)
Sample/Detector Table | Curved bearing guides are over-constrained | Can lead to bearings binding, detector arm cannot be moved | 10 | Bearing sides are not well-aligned | 3 | 1) Machine the bearing mounting surfaces together to ensure the best tolerances 2) Planned inspections of critical dimensions before assembly | 3 | 90 | Could have assembly fixtures to ensure correct alignment… could consider changing the design to use only one bearing rail rather than two
Sample/Detector Table | Curved guides themselves may not provide good geometry for the wire races/balls | Can lead to reduced bearing life, eventual binding & bearing failure which stop detector arm motion | 10 | Fabrication of bearing guides cannot produce the tolerances needed | 2 | 1) Guides will be machined with tight tolerances, but not excessively so. Expect this will be sufficient. | 5 | 100 | Confirm with vendor | Rudy Thermer
Sample/Detector Table | A motor fails | Axis cannot actuate, experiment cannot be set up | 10 | wire failure, connection failure, motor could overheat, lead screw binds | 2 | Components are conservatively designed with large safety margins | 2 | 40 | 1) Could have spare motors on hand (this would be for operations, not for initial installation) 2) Could design a cover to catch spilled sample before it drips onto the motor or lead screw
Hexapod | A motor fails | Sample cannot be correctly aligned | 10 | wire failure, connection failure, motor could overheat, sample spills on motors, user places too much weight on hexapod, hexapod is damaged during transport | 2 | 1) This component is designed for applications like this. Failure is unlikely unless the user damages the hexapod. 2) Generally when the hexapod is used, there will be an oversized mounting plate on top that the sample environment is attached to… acts as a shield for the hexapod motors 3) hexapod is rated to handle 200lbs and all sample environments are expected to be less than that. | 2 | 40 | 1) Label the hexapod with weight limits 2) Create a baseplate to be used as a lifting structure for the hexapod
Sample Handler | A motor fails | Cannot shift between samples without user intervention | 7 | wire failure, connection failure, motor could overheat, lead screw binds, sample spills on motors | 4 | Components are conservatively designed with large safety margins | 2 | 56 | Could have spare motors on hand (this would be for operations, not for initial installation)
Shutter/ Attenuators | Elastomer seals of the air cylinders fail in the radiation environment of the bunker | will not be able to actuate the cylinders, attenuators may be stuck in the beam (or out of the beam) until the bunker can be accessed for maintenance | 10 | Radiation damage | 10 |  | 5 | 500 | 1) Specify EPDM seals (if possible) which are radiation resistant) 2) consider designing radiation shielding for the cylinders 3) test standard cylinders in a radiation enviroment to check for frequency of failures
Shutter/ Attenuators | Air cylinder lubrication fails in the radiation environment of the bunker | will not be able to actuate the cylinders, attenuators may be stuck in the beam (or out of the beam) until the bunker can be accessed for maintenance | 10 | Radiation damage | 10 | None currently | 5 | 500 | 1) Specify EPDM seals (if possible) which are radiation resistant) 2) consider designing radiation shielding for the cylinders 3) test standard cylinders in a radiation enviroment to check for frequency of failures
Slits | Slits fail to actuate | Cannot set up the experiment | 10 | internal lead screws bind (this has happened in JJ Xray slits before) | 3 | JJ Xray knows about the earlier failures and has an updated version of the slits to correct the problem | 2 | 60 | 1) Have spare slit assemblies on hand (at least one of each type). This would be for operations, not for initial operation. 2) May consider switching to SNS-designed slits, one of which has been used for several years without failure
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
Note on F13: Wilson, Danielle: Per John, there is basically no possibility of this happening beyond what we can already adjust for
Note on H13: Wilson, Danielle: Detector will see this during the beamline setup process.
Note on H14: Wilson, Danielle: Many ways to detect this
Note on H15: Wilson, Danielle: Will be immediately obvious if failure exists
Note on H16: Wilson, Danielle: Many ways to detect this
Note on H17: Wilson, Danielle: There are no sensors to detect this failure directly, but table motion sensors will indicate a problem exist.
Note on D18: Wilson, Danielle: Could cause harm to the user
Note on H18: Wilson, Danielle: Currently have no way to detect this
Note on H19: Wilson, Danielle: There are no sensors to detect this failure, but the user will discover it the next time they try to move the slit. Not ideal.
Note on H20: Wilson, Danielle: Currently do not have a way to way to detect if the 3rd slit support arm is damaged
Note on J20: Wilson, Danielle: Not known yet what would be the best corrective action, or if any of these listed would be a solution.
Note on H21: Wilson, Danielle: No sensors to detect this failure directly, but sensors to detect if motion has failed
Note on H23: Wilson, Danielle: May not know this has failed till an end-of-travel move is commanded.
Note on H25: Wilson, Danielle: Would not be able to detect this problem until it caused enough bearing damage to create binding & would not be able to detect it directly, but would be able to detect the motion failure.
Note on H26: Wilson, Danielle: No way to detect this condition until a motion failure occurs. Not ideal.
Note on H27: Wilson, Danielle: Cannot detect a motor failure directly, but can detect a motion failure.
Note on F28: Wilson, Danielle: Connector, wiring, overheating is very rare. Users are very unlikely to overload the hexapod if told not to. Sample spillage is rare, and the hexapod may not be damaged at those times.
Note on H28: Wilson, Danielle: Cannot detect a motor failure directly, but can detect a motion failure.
Note on H29: Wilson, Danielle: Cannot detect a motor failure directly, but can detect a motion failure.
Note on D30: Wilson, Danielle: If attenuators fail out of the beam, it would be a 5 rating… inconvenient, but the user could potentially work around it. If they fail in the beam, they cannot be removed until the proton beam is shut down, and likely would not be enough neutron flux for an experiment.
Note on F30: Wilson, Danielle: Not known what the occurrence will be on this, so making it a 10 to force corrective action
Note on H30: Wilson, Danielle: There will be sensors to detect cylinder motion, which may themselves fail in a radiation environment (unknown)
Note on D31: Wilson, Danielle: If attenuators fail out of the beam, it would be a 5 rating… inconvenient, but the user could potentially work around it. If they fail in the beam, they cannot be removed until the proton beam is shut down, and there likely would not be enough neutron flux for an experiment in the meantime.
Note on F31: Wilson, Danielle: Not known what the occurrence will be on this, so making it a 10 to force corrective action
Note on H31: Wilson, Danielle: There will be sensors to detect cylinder motion, which may themselves fail in a radiation environment (unknown)
Note on D32: Wilson, Danielle: If attenuators fail out of the beam, it would be a 5 rating… inconvenient, but the user could potentially work around it. If they fail in the beam, they cannot be removed until the proton beam is shut down, and there likely would not be enough neutron flux for an experiment in the meantime.

## Severity Scale
Severity Scale
Adapt as appropriate
| Effect | Criteria: Severity of Effect | Ranking
| Very Extreme | User cannot run their experiment within their alloted time window | 10
| Extreme | User might only be able to run <50% of their experiment within their alloted time window, and may only be able to collect <50% of the desired data | 9
| Very High | User may only be able to run >50%, <100% of their experiment within their alloted time window, and may only be able to collect <50% of the desired data | 8
| High | User may only be able to run >50%, <100% of their experiment within their alloted time window, but is able to collect >50% of the desired data | 7
| Moderate | User is able to run all of their experiment and collect >80% of the desired data | 6
| Low | User is able to run all of their experiment and collect 100% of the desired data with <6 hours required to effect repairs or devise a workaround to the failure | 5
| Very Low | User is able to run all of their experiment and collect 100% of the desired data with <4 hours required to effect repairs or devise a workaround to the failure | 4
| Minor | User is able to run all of their experiment and collect 100% of the desired data with <2 hours required to effect repairs or devise a workaround to the failure | 3
| Very Minor | User is able to run all of their experiment and collect 100% of the desired data with <1 hour required to effect repairs or devise a workaround to the failure | 2
| None | No Effect | 1

## Occurrence Scale
Occurrence Scale
| Probability of Failure | Time Period | Per Item Failure Rates | Ranking | Event Frequency Level | Estimated Annual Likelihood of Occurrence | [7 empty] | Description
| Very High | <100% of usage | >= 1 in 2 | 10 |  | Exp | Base | Exp | Exp | Annual |  | Occurs once per
|  | <80% of usage | 1 in 3 | 9 | Anticipated (A) | 10–1 > p > 10-2 | 10 | -1 | -2 | 0.1 | 0.01 | 10 years | 100 years | Incidents that may occur several times during the lifetime of the facility. (Incidents that commonly occur) | SIL 1
| High | <60% of usage | 1 in 8 | 8 | Unlikely (U) | 10-2 > p > 10-4 |  | -2 | -4 | 0.01 | 0.0001 | 100 years | 10,000 years | Events that are not anticipated to occur during the lifetime of the facility. | SIL 2 ~ 4
|  | <50% of usage | 1in 20 | 7 | Extremely Unlikely (EU) | 10-4 > p > 10-6 |  | -4 | -6 | 0.0001 | 1e-06 | 10,000 years | 1,000,000 years | Events that will probably not occur during the life cycle of the facility.
| Moderate | <40% of usage | 1 in 80 | 6 | Beyond Extremely Unlikely (BEU) | 10-6 > p |  | -6 | ∞ | 1e-06 | ∞ | 1,000,000 years | ∞ | Events that are unlikely to occur during the life cycle of the facility.
|  | <30% of usage | 1 in 400 | 5
|  | <20% of usage | 1 in 800 | 4
| Low: | <10% of usage | 1 in 1,500 | 3 | Reference
| Very Low | <5% of usage | 1 in 3,000 | 2
| Remote | 0-2% of usage | 1 in 6000+ | 1

## Detection Scale
Detection Scale
| Detection | Criteria: Likelihood the existence of a defect will be detected by process controls before next or subsequent process, -OR- before exposure to a client | Ranking
| Almost Impossible | No known controls available to detect failure mode | 10
| Very Remote | Very remote likelihood current controls will detect failure mode | 9
| Remote | Remote likelihood current controls will detect failure mode | 8
| Very Low | Very low likelihood current controls will detect failure mode | 7
| Low | Low likelihood current controls will detect failure mode | 6
| Moderate | Moderate likelihood current controls will detect failure mode | 5
| Moderately High | Moderately high likelihood current controls will detect failure mode | 4
| High | High likelihood current controls will detect failure mode | 3
| Very High | Very high likelihood current controls will detect failure mode | 2
| Almost Certain | Current controls almost certain to detect the failure mode. Reliable detection controls are known with similar processes. | 1

## SxO Visualization
Optional Visualization
Severity | 10
9
8
7
6
5
4
3
2
1
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Occurrence