## Slide 1: QIKR Neutron Detection
John F. Ankner
QIKR Instrument Scientist
QIKR Motion Preliminary Design Review
December 16, 2024

## Slide 2: Detecting QIKR
Overview
Measuring reflectivity
  Bring source to sample
  Define beam
  Position sample and detector
  Normalize to direct beam
    Expected count rates
    Revisit requirements
    Mitigation strategies
Summary

## Slide 3: Detecting QIKR
Overview
Measuring reflectivity
  Bring source to sample
  Define beam
  Position sample and detector
  Normalize to direct beam
    Expected count rates
    Revisit requirements
    Mitigation strategies
Summary

## Slide 4: QIKR will make time-resolved reflectometry routine
Source
Bunker
Cave
Control
Hutch
STS flux and wavelength bandwidth will enable routine single-setting measurement of specular reflectivity in seconds to minutes, depending on the sample
“Cinematic” data collection at one instrument setting and event-based data reduction will enable post-experiment tracking of structural evolution in time over a broad scientific landscape
[Density × Avogadro’s number /
Molecular mass]
Monolith
Notes: The picture on the left shows how neutron waves reflect from layers of different refractive index nl below the sample surface
Refractive index depends on the density of scatterers within a molar volume, e.g., monomer or crystal unit cell
The picture on the upper right shows the upper and lower QIKR end stations with cartoon men for scale

## Slide 5: A sampling of scientific communities served by QIKR
Soft Matter
Polymer diffusion
Chemical transformation of reactive films
Hydrogels
Fouling
Structure-properties of films under shear
Synthetic membranes
Responsive films
Polymer brushes under shear
Surface modification
Reactions at oil/water interfaces
Energy Materials
Solid-electrolyte interphase
Organic photovoltaics
Ionic liquids
Corrosion
Mesoporous films
Conjugated polymer films
Metal-harvesting polymers
Biomaterials
Model membranes
Lipid flip-flop
Structure of transmembrane peptides
Biocompatible coatings
Surfactant and phospholipid monolayers
Drug delivery
Protein conformation to membranes
Influence of synthetic nanoparticles on membrane structure
[Courtesy R. Ashkar]
Notes: To be updated with TOCs

## Slide 6: Diverse environments
θi fixed
θs
θd
θ-2θ geometry
θd = 2θs
θ-θ geometry
θd = -θi
θs = 0, fixed
θd
θi
θs = 0, fixed
θd
θi
Two end stations

## Slide 7: QIKR detector requirements
Charge 1

## Slide 8: Detecting QIKR
Overview
Measuring reflectivity
  Bring source to sample
  Define beam
  Position sample and detector
  Normalize to direct beam
    Expected count rates
    Revisit requirements
    Mitigation strategies
Summary

## Slide 9: Neutron reflectivity from Ir/Si standard
t
r
θ
θ
Q = 4π sinθ / λ
NR characterized by wavevector Q, a function of incident angle θ and wavelength λ
No useful information below Qmin
Qmax background limited
=> 0.005 Å-1 < Q < 0.5 Å-1
Fit
Notes: Decibumps?

## Slide 10: Single setting Q range
STS QIKR-L
L = 20.0 m
f = 7.5 Hz
λmin = 2.5 Å
D = 11.5
Wavevector transfer
Q = 4π sinθ / λ
and wave-particle duality
h2 / 2mλ2 = mv2 / 2
determine Q dynamic range
where
h/m = 3956.0 m-Hz-Å
L = instrument length (m)
f = repetition rate (Hz)
λmin = minimum wavelength (Å)
Equivalent fit quality
to full Q range

## Slide 11: Detecting QIKR
Overview
Measuring reflectivity
  Bring source to sample
  Define beam
  Position sample and detector
  Normalize to direct beam
    Expected count rates
    Revisit requirements
    Mitigation strategies
Summary

## Slide 12: Transport source to sample (QIKR-L)
Vertical:
Horizontal:
Double disk
chopper
(8.0 m)
Source
Diverging tapered
from 2.65 to 10 cm high over 11.13 m
Converging tapered
from 10 to 2 cm high over 2.87 m
Guide
begins
(2.0 m)
Guide
ends
(16.0 m)
Straight 3 cm wide
Curved bender
3 channels, 3 cm wide
Sample
Table
(18.0 m)
Slits
Detector
(20.0 m)
Neutron guide filters out unwanted neutrons,
delivering wanted neutrons (3×2 cm2) to the sample

## Slide 13
Guide views – looking upstream
QIKR-U
QIKR-L
Monolith
Bunker
Caves

## Slide 14: Guide system provides multiple beams
Characteristics
Three beams
…separated in space (y)
…and angle γy
Select which to use in θ-θ geometry
to span needed Q range
(QIKR-L McStas simulation)

## Slide 15: Detecting QIKR
Overview
Measuring reflectivity
  Bring source to sample
  Define beam
  Position sample and detector
  Normalize to direct beam
    Expected count rates
    Revisit requirements
    Mitigation strategies
Summary

## Slide 16: Slits define resolution and footprint
lud
su
sd
F
ss = F sinθi
θi
lds
(not to scale)
Which slit settings su and sd will
confine beam within footprint F
at incident angle θ with angular
divergence (δθ/θ)?
What is the integrated acceptance?
=>

## Slide 17: Incident Table – holds slits and optics
Preliminary design of an incident table for QIKR
Translating ball screw enables an accurate pivoting in the range ± 5°. Short support for more stability.
A rigid bearing ensures
precise alignment of the slits.
A platform open to the sides, with sufficient distance to the beam line, offers flexible installation space for future installations.
Adjustable machine feet with vibration absorbing pads.
Manually adjustable slit position with safety brake in z-direction.
Adjustability in the xz plane.
Hollow profiles together with energy chain enable protected cable
installation.
The axis of rotation ideally lies on the face and the symmetry plane of the neutron guide.
[courtesy Rudy Thermer]

## Slide 18: Footprint calculator – maximum slits
Resolution
0.02 < δθ/θ < 0.07
implies
0.1 mm < s < 10 mm

## Slide 19: Footprint calculator – minimum slits
Beam Size
and Footprint
Smallest: 0.1 mm
δθ/θ = 0.02 @ θ = 0.5°
=> δθ = 0.01°
[Req. S.04.08.04-R2]

## Slide 20: Detecting QIKR
Overview
Measuring reflectivity
  Bring source to sample
  Define beam
  Position sample and detector
  Normalize to direct beam
    Expected count rates
    Revisit requirements
    Mitigation strategies
Summary

## Slide 21: Sample Table and Detector Arm unit
Available with different load capacities i.e., various spindle sizes, gear ratios and accuracies.
PI H-850 Hexapod
Payload: 250 kg
FDet
FLift
15°
Preloaded linear ball bushings
ZIMM Screw Jack
With an ideal load-bearing arrangement in relation to the spindle lifting nut.
Increased stiffness
due to a free length
Reduction of 50%.
THK-TU-Series Slide table
providing high load capacity
and accuracy in a single
Structure.
The design employs well-developed standard components (increased reliability, reduced costs).
Nadella GLK (NIRO) Spherical-plain bearing Radial clearance: 0 -10 µm!
Compensatory force transmitting
link.
Hexapods offer six degrees of freedom of movement in the most compact space.
Robust guide system
∅ 60 mm
Yaw swivel unit
Spindle nut
THK Curved linear guides
[courtesy Rudy Thermer]

## Slide 22: Sample Table and Detector Arm unit
Available with different load capacities i.e., various spindle sizes, gear ratios and accuracies.
PI H-850 Hexapod
Payload: 250 kg
FDet
FLift
15°
Preloaded linear ball bushings
ZIMM Screw Jack
With an ideal load-bearing arrangement in relation to the spindle lifting nut.
Increased stiffness
due to a free length
Reduction of 50%.
THK-TU-Series Slide table
providing high load capacity
and accuracy in a single
Structure.
The design employs well-developed standard components (increased reliability, reduced costs).
Nadella GLK (NIRO) Spherical-plain bearing Radial clearance: 0 -10 µm!
Compensatory force transmitting
link.
Hexapods offer six degrees of freedom of movement in the most compact space.
Robust guide system
∅ 60 mm
Yaw swivel unit
Spindle nut
THK Curved linear guides
[courtesy Rudy Thermer]
Positioning
Axes of ST and DA
must coincide in beam
DA rotates about
sample

## Slide 23: Detecting QIKR
Overview
Measuring reflectivity
  Bring source to sample
  Define beam
  Position sample and detector
  Normalize to direct beam
    Expected count rates
    Revisit requirements
    Mitigation strategies
Summary

## Slide 24: Measuring specular reflectivity
Divide reflected by incident intensity
Carefully measure primary data . Three options for measuring :
1. Measure via series of attenuated incident apertures
2. Measure via pixelated scan of incident aperture
3. Measure via reflectivity of known reference sample

## Slide 25: Simulate QIKR direct-beam rates (Id)
Maximal and typical beams were simulated using McStas
Large Id at short λ
Three-decade Id dynamic range
To measure and use Id
Create tractable count rate
Even out counting statistics in λ

## Slide 26: Solution: normalize to reflectivity standard
Reflect from D2O and measure
Typical configuration’s rate reduced to ~20 kHz
Dynamic range flattened
Charge 1

## Slide 27: Solution: normalize to reflectivity standard
Reflect from D2O and measure
Maximal configuration’s rate also reduced to <~20 kHz – steeper incident angle lowers D2O reflectivity
Dynamic range flattened
Charge 1

## Slide 28: Challenge: poor short-wavelength performance
Reflect from D2O and measure
Maximal configuration’s rate also reduced to <~20 kHz – steeper incident angle lowers D2O reflectivity
Dynamic range flattened
Issue: D2O reflectivity too low at short λ
For λ < 4 Å, RD2O small and difficult to measure, yielding low rate
  Develop a different standard e.g., 20 Å Ni / SiO2
  Or measure short λ via attenuation and/or pixelation

## Slide 29: What about reflected beam (IR)?
Reflect from Ni / SiO2
Maximal configuration
  θ = 0.83°; δθ/θ = 0.07
  IR > 106 n sec-1
What to do?

## Slide 30: What about reflected beam (IR)?
Reflect from Ni / SiO2
Maximal configuration
  θ = 0.83°; δθ/θ = 0.07
  IR > 106 n sec-1
What to do
Use a steeper angle
  θ = 2.5°; δθ/θ = 0.07
  IR < 105 n sec-1
Charge 1

## Slide 31: What about reflected beam (IR)?
Reflect from Ni / SiO2
Maximal configuration
  θ = 0.83°; δθ/θ = 0.07
  IR > 106 n sec-1
What to do if you must see R=1
Use a steeper angle
  θ = 1.5°; δθ/θ = 0.07
  IR > 105 n sec-1
  Dead-time correct?
  Attenuate?
Charge 1

## Slide 32: Effect of attenuator
“Achromatic” attenuators
Use e.g., polycarbonate
(PC) or Plexiglas (PMMA)
Weakly λ dependent – alters spectrum
Multiple measurements needed to correct spectral distortion
Thick attenuator behaves like a moderator

## Slide 33: Detecting QIKR
Overview
Measuring reflectivity
  Bring source to sample
  Define beam
  Position sample and detector
  Normalize to direct beam
    Expected count rates
    Revisit requirements
    Mitigation strategies
Summary

## Slide 34: Reflectometer detector options
MWPC
Ex. instruments:
LiqRef, MagRef, SupeADAM
(also on SANS, diff.)
Built by:
(BNL), Denex, ILL
Advantages:
High efficiency
Low gamma
Limitations:
Disadvantages:
(Very) low rate
Thick window
Multi-Tube
Ex. instruments:
FIGARO, D17, PLATYPUS
(also on SANS, diff.)
Built by:
ILL
Advantages:
High efficiency
Low gamma
Limitations:
Asymmetric pixels
Disadvantages:
Medium rate
WLSF
Ex. instruments:
INTER
(also on diffraction)
Built by:
ISIS, SNS
Advantages:
High efficiency
Low gamma
Limitations:
Asymmetric pixels
Disadvantages:
Medium rate
Multi-Blade
Ex. instruments:
None
Planned: ESTIA, FREIA
Built by:
ESS
Advantages:
Medium-high rate
Low gamma
Limitations:
Asymmetric pixels
~40% efficiency
ASIC readout
Disadvantages:
Uniformity
Ex. instruments:
None
(several diffractometers)
Built by:
SNS
Advantages:
Medium-high rate
Sufficient resolution
Limitations:
Gamma sensitivity
Rate currently not enough
Disadvantages:
Low radiation hardness
Ex. instruments:
None
Built by:
SNS (pending R&D)
Advantages:
Very high rate
Limitations:
Large Pixels
Disadvantages:
High gamma sensitivity
Pixel gaps
Ex. instruments:
Imaging at LANL, PSI
(Tested in many instruments)
Built by:
LoscoVision / ASI
Advantages:
High rate
Very high resolution
Flexible configuration
Likely several instruments
Limitations:
Improvements pending
Disadvantages:
Anger Camera
Pixel Camera
LumaCam

## Slide 35: Summary
The specifications for the QIKR detector have been derived from extrapolation of existing FTS-LR capabilities
No detector will be able to measure the anticipated direct beams directly
The expected instantaneous rate for reflectivity measurements suggests increasing specification S.04.08.05-R2 from 2×104 to 105
QIKR has been designed to enable three different mitigation strategies
  Attenuated direct beams
  Pixelated direct beams
  Normalization to reflectivity
Dead-time correction will be applied as needed
The appropriate strategy/strategies will be chosen based on detector capability
The better the detector, the easier the mitigation
Charge 1

## Slide 36: Thanks
Anton Khaplanov
Danielle Wilson
Ryan Butz

## Slide 37: Compile into table of direct beam “runs”
Direct beams in counts/mC – no scale factors required

## Slide 38: Et voilà: specular reflectivity normalized to

## Slide 39: Location of QIKR
QIKR will sit at port ST02 and view the cylinder moderator
QIKR
Beamline
ST02
STS Beamlines in the South Instrument Hall

## Slide 40: QIKR-L instrument cave
Detector Arm
Incident Table
Guide
Sample Table
Shield
Walls
Cave Entrance
[Requirements A,B,F]

## Slide 41: Attenuated apertures
Start with measurable spectrum
Place attenuator, re-measure
Repeat as needed

## Slide 42: In practice – constant θ = 0.6°
Measure at reduced apertures
(<1 kHz for Brookhaven 3He)
Determine scale factors by a series
of attenuated measurements
Fj = Aj + Bj λ
Apply factors to determine
actual spectrum

## Slide 43: In practice – constant 2.5 < λ < 6 Å
Measure at reduced apertures
(<1 kHz for Brookhaven 3He)
Determine scale factors by a series
of attenuated measurements
Fj = Aj + Bj λ
Apply factors to determine
actual spectrum – gets
softer as slits open wider