## Slide 1: QIKR Detector PDR: Mechanical Overview
D. Wilson, Lead Engineer
December 16, 2024

## Slide 2: Outline
QIKR Overview
  Review Scope
  Documentation
  High level QIKR description
Detector & Beam Monitor
  Beam Monitor locations
  Detector details

## Slide 3: QIKR Overview

## Slide 4: Review Scope
Included
QIKR Detector
QIKR Beam Monitor (Overview)
Not Included
Details of QIKR Beam Monitor design… this is a design that is copied from ILL

## Slide 5: Requirements, CQLD, Interface, FMEA Documents
Requirements Documents:
  QIKR Instrument Requirements: S04080100-SRD10000-R02 (not released)
    The latest revision (R02) has not been released yet, awaiting optics alignment updates
CQLD Detector:
  Quality Level 2, Configuration Level 3: S04080500-QAI10000-R00 (released)
    Quality Level 2 is primarily because of the impact to the user of a detector failure on QIKR (there is only one detector per beamline)
CQLD Beam Monitor:
  Quality Level 3, Configuration Level 3: S04080500-QAI10001-R00 (released)
Interface Documents:
  Interface to ICS (detectors): S01020500-IST10117-R00 (released)
FMEA for Detector: S04080500-FMA10000-R00 (not released)
  Created by D.Wilson, J.Ankner, A.Khaplanov
  One item of concern (discussed in later slides)

## Slide 6: Interface Details, ICS to QIKR Instrument Detector

## Slide 7: STS Instrument Suite
ST01
EXPANSE
ST02
QIKR
ST03
CUPI2D
ST04
BWAVES
ST05
CHESS
ST06
Future
ST07
VERDI
ST08
PIONEER
ST09
CENTAUR
Monolith Region
South Hall
(ST01-ST09)
Target Bay
North Hall
(ST10-ST18)
Proton Beam

## Slide 8: QIKR – Schematic View
The shutters also contain the beam monitors
The End Station consists of the Incident Table with Slits and Frame Overlap Mirror, the Sample Table, the Detector Table, and the Sample Changer
As oriented in this view, QIKR-A occupies the back half of the instrument wedge & cave, and QIKR-B occupies the front half
QIKR-A End Station
QIKR-B End Station
QIKR-A
QIKR-B
QIKR-A
24m moderator-to-detector distance
20m moderator-to-detector distance

## Slide 9: QIKR –3D Model View
Neutrons
Maintenance Shield
Shutters, Beam Monitors
Incident Table with Slits & Frame Overlap Mirror
Sample and Detector Tables, Sample Changer
QIKR-A
QIKR-B
User stands on opposite sides of the end station, upper vs. lower
End Station

## Slide 10: Installation of Components
2-ton monorail in each cave for component installation, & sample environment changes
All QIKR in-bunker components will be installed by removing the bunker roof panels and using the 50-ton high bay gantry crane unless they are light enough to be hand-carried
Most QIKR in-cave components will be installed by bringing the components into each cave on a cart, then carrying them to their final location with a 2-ton monorail hoist
The monorail will also be used during operations to swap sample environments and to service end station components
Note: The detector and its associated components have a low enough weight to allow manual assembly without the monorail
Beam stops are hinged and swing out of the way to allow user or component access when needed

## Slide 11: Cable Routing
QIKR-B patch panel (on inner wall)
QIKR-B power panel (on outer wall)
QIKR-B additional routing to the roof
QIKR-B motion & detector cable routing to the roof & control racks
QIKR-A motion & detector cable routing to the roof & control racks
QIKR-A additional routing to the roof
IPPS boxes for Shutters (on exterior bunker wall)
QIKR-B routing through bunker wall pass-thru & roof for in-bunker components
Control racks for both QIKR-A and QIKR-B
Cable routing to cave both QIKR-A and QIKR-B
All IPPS cables
QIKR-B’s motion and detector cables are routed through QIKR-A’s cave, which allows the cables to be under 30m in length
QIKR-A’s motion and detector cables are under 30m in length as well (routed directly up through the roof)
QIKR-A power panel (on inner wall)
QIKR-A patch panel (on inner wall)

## Slide 12: Beam Monitor

## Slide 13: Shutter
The Shutter blocks the neutron beam to allow personnel access to the cave… there is one shutter per beamline (two total)
The QIKR Shutter operation is different from standard in two ways:
  The “open” position also contains the attenuators
  There is a third position for a beam monitor
The need to minimize guide gaps and windows drives the multi-functionality of the shutter
Upper beamline (QIKR-A) Shutter
Lower beamline (QIKR-B) Shutter
Blocked position
Open position (with optional attenuation)
Beam Monitor position
Downstream side (looking upstream)
Design credit: Rudy Thermer

## Slide 14: Shutter Positions
A) Beam blocker in position
B) Open position with attenuator
C) Multitube beam monitor in position
Neutrons
Neutrons
Neutrons
Design credit: Rudy Thermer

## Slide 15: Shutter Components
Downstream side (looking upstream)
PPS “Beam Monitor”
PPS “Beam Open”
PPS “Beam Blocked”
2-stage air cylinder, fail-safe for electrical failure
Fabricated aluminum plates
Beam Blocker
10cm tungsten +
1cm B4C
Aluminum extrusion frame (similar to standard design)
Beam Monitor
Bearing rail (stainless steel)
Attenuator assembly
PPS flag
Beam Monitor Attachment
Fabricated carrier for the beam blocker
Design credit: Rudy Thermer
Elongated bracket holes allow adjustment in x-direction
Nested shoulder screw acts as a guide pin

## Slide 16: Component Accessibility for Assembly & Disassembly
Multitube monitor: accessible from below.
Multitube monitor & Attenuator: accessible from above the guide
Design credit: Rudy Thermer

## Slide 17: Beam Monitor Requirements
There are no specific beam monitor requirements for QIKR

## Slide 18: Acquisition strategy
The parts will be effectively COTS, ordered from ILL
Competitive bid not possible… ILL will supply the monitors
Monitor is not available from a US distributor
No expected supplier constraints (unless ILL closes before STS orders parts)
No exotic or hard-to-obtain materials

## Slide 19: Manufacturing strategy
There are no fabricated parts
Assembly of the beam monitor into the beamline will be done in-house
  No assembly needed at the monitor level itself

## Slide 20: Installation Plan
Beam Monitor will be placed manually onto the shutter & tested for basic functionality prior to being installed in the beamline
Riggers to place the shutter assembly into the bunker with the 50-ton overhead high bay crane (through the bunker roof)
  Installation should occur after the bunker wall and monolith guide sections are installed, but before the remaining in-bunker guide sections are installed
SNS technicians & engineers to do final adjustments, connect cabling, test component function while connected to bunker air & instrument power, PPS, ICS

## Slide 21: Beam Monitor Costs & Labor Hours
Beam Monitor | Mat’l Cost | Labor Cost | Labor Hours
Procure (P6) | $46,489 | $5,836 | 28
Install (P6) | $1,954 | $8,570 | 50
P6 vs. Current | Labor Type | Procure/Asm | Install
Source = P6 | scientist | 20 | 20
 | engineer |  | 6
 | designer |  | 
 | technician | 8 | 24
 | Davis-Bacon |  | 
Material & Labor Costs
Labor Hours
ICS has additional hours for install, not funded from the Instrument budget
Only the QIKR-B monitor will be purchased and installed

## Slide 22: Detector

## Slide 23: End Station – Overview
Upper beamline (QIKR-A)
Lower beamline (QIKR-B)
Incident Table Assembly (upper and lower versions)
Sample and
Detector Table Assembly
End station components:
  Incident Table
  Sample/Detector Table
  Beam Conditioning Slits
  Frame overlap mirror (FOM)
  Hexapod
  Sample Changer
  Detector & Shielding
The user stands on the opposite side of the upper end station vs. the lower one
  Some features on the detector may be mirrored
User stands here
Design credit: Rudy Thermer

## Slide 24: Incident Table – Operation Overview
The incident table’s purpose is to select one of the 3 main neutron beams exiting the guide
A beam is selected by moving the slits and Frame Overlap Mirror (FOM) as a unit to be in line with the desired beam
  The slits allowed the selected beam to pass through while blocking the others
This detector must look at one of three possible sample locations each separated vertically by ~60mm
At the guide exit, the neutron beam splits into 3 main components separated by 1.5°
Guide
Neutrons
Incident Table pivot point
Frame Overlap Mirror (FOM)
Slits
Tie-downs to the cave floor (on all 4 sides)
Design credit: Rudy Thermer

## Slide 25: Sample/Detector Table
Hexapod (electric motors) with spacers for fine positioning of the sample
This sample table moves up/down to place the sample in the path of one of the three beams
The detector arms rotate about the sample center location (+15°) and are carried up/down vertically with the sample table
Curved wire-race bearings (on both sides) for the detector arm rotation
Sample table and detector arms can rotate 15° about a vertical axis through the sample center
Detector, downstream slit, and detector snout (shields the detector from background) all sit on a table mounted to the detector arms via six-strut linkages to allow initial alignment (manual)
Detector & shielding will change slightly as the detector geometry is finalized, but will be very similar to what is shown
The detector and sample table move vertically together (motorized)… there is no separate vertical motion for those components
  Reduces axes & cost
The detector and sample table also rotate together (motorized) about a vertical axis through the sample nominal center
Fine sample positioning is handled by an electric motor hexapod sitting on a stack of shims
  The shims provide larger height adjustments for tall vs. short sample environments
The detector rotates about an x-axis (motorized) through the sample center using curved guide rails
  The detector is always looking at the sample center, even if the sample table changes height
Base plate is bolted to a levelled grout plate (not shown) in the cave floor
Design credit: Rudy Thermer

## Slide 26: Detector Platform
The detector platform is attached to the detector arm via 6-strut linkages… 6 turnbuckles that fully constrain and allow position adjustment of all DOF’s
  Struts allow ~25mm (1”) of linear adjustment
The detector platform has lift features for installation via the overhead monorail crane
Design is a placeholder at the moment… will be completed when the detector and snout design is finished (final detector geometry TBD)
There will be mounting features for the detector, the snout, and the slit that allow some relative position adjustment between the 3 components
6-strut linkages… allow linear and angular adjustment of the detector, snout, and slit as a unit (adjustment range ~25mm)
Fabricated aluminum plates, screwed together (screws not shown)
Fabricated aluminum or steel brackets to attach the 6-struts
Lift features for installation
Attach points for the detector, snout, & slit… final design TBD pending detector design finalization
Design credit: Art Diaz
Slit and shield plate (orange)
The components will be mounted to an optical rail (not shown) once the detector design is finalized

## Slide 27: Motions of Detector Arm and Sample Table
The 3 neutron beams for QIKR-B
Sample positions needed for each beam
Detector
Detector Arm
Sample Table
Sample/ Detector Table Base
Hexapod
Sample Table Vertical Travel
Detector Arm rotates about a horizontal ‘X’ axis drawn through the nominal sample center
Each sample position separated by ~60mm
Sample Table moves vertically to place the Detector Arm’s center of rotation on the desired sample position
Sample Table also rotates toward the user about a vertical ‘Y’ axis drawn through the nominal sample center
Vertical ‘Y’ axis through nominal sample center

## Slide 28: Detector Description
Mechanical Details of Detector
  The detector consist of:
    a light-tight box that has a scintillator coating on the interior of the surface that faces incoming neutrons
    a 45° mirror that reflects visible light into a camera lens
    a camera lens that focuses the incoming light (lenses can be changed as needed for different fields of view)
    an intensifier that multiplies the incoming photons
    a camera that captures the visible light and relays an image
    a motor with encoder that moves the camera, lens, and intensifier in order to adjust the focus
Detector geometry not final yet
  Needs to be up-sized to give QIKR’s required 20cm x 20cm detection area… details of that have not been finalized (must work with Losko Vision)
  Perhaps use a two-camera design instead of only one?

## Slide 29: LumaCam Detector
Light-sensing camera
Focusing motor
Bellows to allow focusing motion while maintaining light-tightness (camera lens moves inside the bellows)
Scintillator coats the interior of this aluminum housing surface
45° mirror to reflect the scintillation light into the camera lens
Light intensifier
Light-tight box (lid not shown)
Motor + encoder cable

## Slide 30: LumaCam Detector
Fixed neutron-sensing portion of detector
Movable camera-focusing portion of detector (camera, lens, intensifier, one end of the bellows)
Camera output data cable (fiber-optic)
Camera power cable
Image intensifier
Camera timing signal cable (not shown)
Low voltage cable to control amount of signal intensification (volume control)
5V TTL signal cable
Low voltage intensifier power cable (not shown)

## Slide 31: LumaCam Detector
Neutrons
Neutrons
Light
Camera lens

## Slide 32: Detector Shielding Description
There is shielding around the detector to prevent unwanted neutrons from getting to the scintillator
  A rectangular pyramid “snout” extends between the detector and the slit downstream of the sample
  A shield box encloses the scintillator housing on all sides
  A shield plate is assembled on the upstream side of the slit
The camera is not expected to need full shielding & should remain un-enclosed to avoid potential overheating issues

## Slide 33: Detector Shielding Details
Shielding is comprised of multi-layer panels
  The thickest layer is borated polyethylene (PE)
  The interior layer is B4C (glued on)
  The exterior layer is aluminum sheet (glued on)… the PE must be covered per fire protection rules
The PE layer contains threaded inserts that are used to screw the panels together
Most screws have thumb-screw caps attached for easy installation/ removal
  Allows for easy access to the interior of the scintillator housing if lens changes are needed
Larger panels have handles, also for easy installation/ removal
Weight of assembly shown = 195kg Weight allowed = 306kg
Aluminum
(blue)
Borated polyethylene (PE)
(green)
B4C (red)
Handles
Screws with thumb-screw caps
Attachment to the detector platform
“Snout”
Camera shielding
Scintillator housing shielding
Design credit: Ryan Butz
Detector platform
Snout supports… will likely attach these to an optical rail

## Slide 34: Detector Shielding – “Snout”
Bottom snout panel is placed first
Side panels are screw-attached to the bottom panel
Top panel is placed and screw-attached to the sides
Weight of each panel is ~33lbs
Snout will eventually be placed on an optical rail that will allow it slide toward and away from the detector to close any gaps in shielding
  Snout is not screw-attached to the scintillator housing shielding, only to the detector platform
Scintillator housing shielding
Camera shielding

## Slide 35: Detector Shielding – Camera & Scintillator
Detector frame screw-attaches to the detector platform
  An 8.6lb rectangular piece of shielding fits beneath… free-floating, not screw-attached… held in place by the detector frame
Two Free-floating panels fit around intensifier tube
  Each panel weighs ~0.6lbs
Free-floating camera shielding sits on top of the panels around the intensifier tube… can attach this to neighboring shielding if needed
  Side panels are each ~6lbs, the top panel is ~2.6lbs
Shield panels for the scintillator housing screw-attach to each other and to the detector platform
  Total weight ~86lbs, heaviest single component weight ~22lbs
Assembled around intensifier tube… free-floating (intensifier not shown)
Camera shielding,3 pieces, sits on shielding below… free-floating (shown as a cross-section)
Detector frame with shielding tucked beneath
Shielding fully encloses scintillator housing (shown as a cross-section here). Panels are screw-attached to each other

## Slide 36: Detector Cable Routing – Description
There are six cables plus a motor power & encoder cable
  Three cables for the camera (signal, power, timing), and three for the intensifier (5V TTL, power, amplification)
  These will all be routed along the detector arm on the side farthest from the user, then drop down to the floor and route to the control racks from there
  Motor power will be kept separate from the remaining cables
No intermediate junction boxes (power supplies, signal conditioners, etc.) are needed
Detector is not expected to need separate grounding, but should have the capability to be connected to ground within the beamline

## Slide 37: Detector Cable Routing – Details
Motor & encoder cable
Camera & intensifier cables
This side of the arm is opposite the user. Cables are routed to the sample location, then run down into the cable channel in the floor.
Sufficient cable slack will be provided here to accommodate the detector arm and sample table table rotations
Motor and encoder cable routes separately from the camera & intensifier cables

## Slide 38: Detector Requirements
These have been achieved by the motion systems

## Slide 39: Detector Concerns from FMEA
Neutrons
Light
Camera lens
Low voltage intensifier power cable (not shown)
User may want to open the light-tight box to change the lens
A reasonable use-case is for the lens to be swapped out for one with a different focal length
Opening the lid of the light-tight scintillator housing to access the lens can cause damage to the intensifier if the intensifier is not powered off first
  there are no controls in place to prevent this & it will be easy to forget the power off step
Recommendation from the FMEA is to design the lid and intensifier power cable routing such that the lid cannot be removed without unplugging the intensifier power

## Slide 40: Acquisition strategy
Detector is effectively a COTS part, customized as needed by LoskoVision
  Shielding is fabricated
Competitive bid not possible on the detector, is possible with the shielding
Shielding components are available from US distributors
No expected supplier constraints (unless LoskoVision closes before STS orders parts)
No exotic or hard-to-obtain materials

## Slide 41: Manufacturing strategy
The fabricated parts will be build-to-print
Detector will be built to a specification
Assembly of shielding will be done in-house, as will some small amount of assembly for the detector

## Slide 42: Installation Plan
The detector platform will be installed using the cave overhead monorail hoist
The detector and shielding will be placed on the platform manually, then will be aligned to the sample location using the six-strut linkages
SNS technicians & engineers will do final adjustments, connect cabling, test component function while connected to bunker air & instrument power, PPS, ICS

## Slide 43: Detector Costs & Labor Hours
Shutter | COTS Mat’l Cost | Labor Cost | Labor Hours
P6 | $1,134,522 | $166,638 | 1215
Current Design | $544,236 | $168,462 | 1231
Delta | -$590,286 | +$1,824 | +16
P6 vs. Current | Labor Type | Procure/Asm | Install
P6 | scientist | 200 | 40
 | engineer | 100 | 15
 | designer |  | 
 | technician | 820 | 40
 | Davis-Bacon |  | 
Current Design | scientist | 200 | 40
 | engineer | 100 | 15
 | technician | 820 | 56
 | Davis-Bacon |  | 
Costs
Labor Hours
ICS has additional hours for install, not funded from the Instrument budget
Current Design labor hours include an additional 16 technician hours for detector shielding assembly
*details of the shielding and detector costs for the current design are shown on slides 46-47
Includes detector and shielding costs for a single camera design
For a 2-camera design, this cost would go to $864,663 for the camera asm plus ~$30,000 additional for the shielding

## Slide 44: Summary
Component weights are within the capacity of the motion components to support them
Cost is within the P6 budget, even if a 2-camera solution is used
Detector and beam monitor both fit well in the space allotted to them
  There is room to grow the detector should the decision be to go with dual cameras in the detector
There is some concern with acquiring the beam monitor, but this is not a critical component for QIKR
The detector is single-sourced with LoskoVision, but ORNL is allowed to produce cameras on its own (if needed) once the contract with LoskoVision is established

## Slide 45: Backup Slides

## Slide 46: Detector Cost Details
Cost estimates in 2024 dollars:
  $133,000 TimePix3 camera double this for 2-camera design
  $50,000 Intensifier double this for 2-camera design
  $45,000 Enclosure + motor + controller + cables + mirror
  $4,000 Scintillator double this for 2-camera design
  $2,000 Lens double this for 2-camera design
  $50,000 Design Costs for Enclosure (NRE)
  $284,000 TOTAL (2024) $473,000 for 2-camera design
Cost estimates in 2031 dollars (last year of procurement)
  $519,163 Total 1-camera design costs
  $26,897 Shielding including additional labor (see next slide)
  $546,060 TOTAL (2031)

## Slide 47: Shielding Cost Details
Aluminum plates are priced at $6.50/lb in 2024 dollars, $8/lb in 2031 using 3% escalation per year
  Assume a $500 charge per plate for machining
  Total costs = $7,454
5% Borated Polyethylene blocks cost $6,945 (tax and delivery included) in 2024 dollars, $8540 in 2031. Converted that into cost/volume to calculate PE shielding cost for each piece
  Assume $500 charge per piece for machining
  Total costs = $15,354
5mm thick Mirrobor, 0.5m x 1m was quoted at $1,060/m2 in 2022, estimated to cost $1,510/m2 in 2031
  Assume 2 man-days of technician hours to cut and assemble the Mirrobor (16hrs @ $114/hr = $1,824)
  Total area of Mirrobor needed = 1,292,960mm2 = 3 panels of 0.5m x 1m Mirrobor needed ($2,265)
  Total costs $4,089