## Slide 1: QIKR Instrument Overview
John F. Ankner
QIKR Instrument Scientist
QIKR Shielding Preliminary Design Review
March 3, 2025

## Slide 2: Outline
QIKR, Briefly
Unwanted Neutrons
  Lost in monolith and bunker
  Lost in instrument cave
Summary

## Slide 3: Outline
QIKR, Briefly
Unwanted Neutrons
  Lost in monolith and bunker
  Lost in instrument cave
Summary

## Slide 4: Location of QIKR
QIKR will sit at port ST02 and view the cylinder moderator
QIKR
Beamline
ST02
STS Beamlines in the South Instrument Hall

## Slide 5: QIKR Instrument Sections
QIKR-U
QIKR-L
Bunker
Cave
Hutch
Lintel
ST02

## Slide 6: QIKR features two independent beamlines (schematic)
The maintenance shield does not perform a beam blocking function
The End Stations consist of an Incident Table with Slits, a Sample Table, and a Detector Arm
As oriented in this view, QIKR-U occupies the back portion of the instrument wedge & cave, and QIKR-L occupies the front portion
QIKR-U End Station
QIKR-L End Station
QIKR-U
QIKR-L
QIKR-U
ST02

## Slide 7: QIKR will make time-resolved reflectometry routine
Source
Bunker
Cave
Control
Hutch
STS flux and wavelength bandwidth will enable routine single-setting measurement of specular reflectivity in seconds to minutes, depending on the sample
“Cinematic” data collection at one instrument setting and event-based data reduction will enable post-experiment tracking of structural evolution in time over a broad scientific landscape
[Density × Avogadro’s number /
Molecular mass]
Monolith
Incident
beam
Reflected
beam
Notes: The picture on the left shows how neutron waves reflect from layers of different refractive index nl below the sample surface
Refractive index depends on the density of scatterers within a molar volume, e.g., monomer or crystal unit cell
The picture on the upper right shows the upper and lower QIKR end stations with cartoon men for scale

## Slide 8: A sampling of scientific communities served by QIKR
Soft Matter
Polymer diffusion
Chemical transformation of reactive films
Hydrogels
Fouling
Structure-properties of films under shear
Synthetic membranes
Responsive films
Polymer brushes under shear
Surface modification
Reactions at oil/water interfaces
Energy Materials
Solid-electrolyte interphase
Organic photovoltaics
Ionic liquids
Corrosion
Mesoporous films
Conjugated polymer films
Metal-harvesting polymers
Biomaterials
Model membranes
Lipid flip-flop
Structure of transmembrane peptides
Biocompatible coatings
Surfactant and phospholipid monolayers
Drug delivery
Protein conformation to membranes
Influence of synthetic nanoparticles on membrane structure
[Courtesy R. Ashkar]
Notes: To be updated with TOCs

## Slide 9: Nano-scale layers deposited on flat substrates
Spin Coating
Layer-by-Layer Growth
Rheometry
Langmuir-Blodgett
deposition
Langmuir Films
Electrochemistry

## Slide 10: Diverse geometries
θi fixed
θs
θd
θ-2θ geometry
θd = 2θs
θ-θ geometry
θd = -θi
θs = 0, fixed
θd
θi
θs = 0, fixed
θd
θi
Two end stations

## Slide 11: Outline
QIKR, Briefly
Unwanted Neutrons
  Lost in monolith and bunker
  Lost in instrument cave
Summary

## Slide 12
Guide views – looking upstream
QIKR-U
QIKR-L
Monolith
Bunker
Caves
Begin
L = 2 m
End
L = 20 m
End
L = 16 m
Guides taper, curve, and are angled up/down, left/right

## Slide 13: QIKR Overview – Guide Path
3-channel Ballistic Bender (BB)
Tapered Guide (TG)
QIKR-U guide
QIKR-L guide
Gaps between Elements
Ballistic Bender blades
View toward the moderator along the z-axis
Three sections in each guide path: Ballistic Bender (BB), Ballistic Guide (BG), and Tapered Guide (TG)
Each guide Element is straight, even in the BB section
In the BB section, each guide element is angled relative to its neighbor to approximate a 156-m radius circle
  The initial 12 elements in each bender section are multichannel, the number of channels = 3
The upper guide initially angles up by 2.5° and rotated by 0.7° toward ST01
The lower guide angles down by 2.5° and rotated by 0.5° toward ST03
Guide Element
Ballistic Guide (BG)
3-channel Ballistic Bender (BB)
Ballistic Guide (BG)
Tapered Guide (TG)

## Slide 14: Guide representation in MCNP
Neutron guide transport calculated using a common STS source and dimensions and coordinates transferred from the engineering model (Creo->IGES) to both McStas and MCNP agreed within 15%

## Slide 15: QIKR-L guide cartoon
Vertical:
Horizontal:
Double disk
chopper
(8.0 m)
Source
Diverging tapered
from 2.65 to 10 cm high over 11.13 m
Converging tapered
from 10 to 2 cm high over 2.87 m
Guide
begins
(2.0 m)
Guide
ends
(16.0 m)
Straight 3 cm wide
Curved bender
3 channels, 3 cm wide
Sample
Table
(18.0 m)
Slits
Detector
(20.0 m)
The QIKR-L and QIKR-U guides have been optimized
to fulfill requirements at minimal cost in supermirror coating

## Slide 16
Neutrons lost in monolith and bunker
Segment
0 (entrance)
1,2 (shield)
3 (chopper)
4 (shutter)
5 (ballistic end)
6 (exit)
A Tale of Loss
Bulk of fast neutrons shed in Ballistic Bender (BB – Segments 1-2)
Little loss in Ballistic Guide (BG – Segments 3-5)
Tapered Guide (TG) trims unwanted thermal and cold neutrons (Segment 6)
→ Shielding must mop up unused neutrons

## Slide 17: Outline
QIKR, Briefly
Unwanted Neutrons
  Lost in monolith and bunker
  Lost in instrument cave
Summary

## Slide 18: Guide system provides multiple beams
Characteristics
Three beams
…separated in space (y)
…and angle γy
Select which to use in θ-θ geometry
to span needed Q range
(QIKR-L McStas simulation)
Guide
Incident Table rotates to align with chosen beam exiting the guide. Slits absorb the unused beams and shape the chosen beam.
slits
QIKR-L

## Slide 19: QIKR end stations: beam selection and detector motion
Upper station (QIKR-U)
Lower station (QIKR-L)
Incident Table Assembly (upper and lower versions)
Sample Table and
Detector Arm Assembly
The 3 available beams for QIKR-L
Sample positions needed for each beam
Detector
Detector Arm
Sample Table
Sample/ Detector Table Base
Hexapod
Detector Arm rotates about a horizontal x axis (out of the screen) passing through the sample center
Hexapod for tilting the sample

## Slide 20: Slits define angular resolution and beam footprint
lud
su
sd
F
ss = F sinθi
θi
lds
(not to scale)
Slits su and sd confine beam within footprint F at incident angle θ with angular divergence (δθ/θ)
→ All unwanted neutrons must be absorbed by shielding (slit blades and housings)
=>
Useful data cannot be collected in the absence of collimating slits, so after commissioning they will remain in place

## Slide 21: Operate in θ-2θ and θ-θ geometries
θi fixed
θs
θd
θ-2θ geometry
θd = 2θs
θ-θ geometry
θd = -θi
θs = 0, fixed
θd
θi
θs = 0, fixed
θd
θi
θ-2θ geometry
Incident beam fixed
Sample moves
Detector moves
θ-θ geometry
Incident beam varies
Sample fixed
Detector moves

## Slide 22: Footprint calculator – maximum slits
Resolution
0.02 < δθ/θ < 0.07
implies
0.1 mm < s < 10 mm
→ Maximum usable slit aperture is smaller than the fully open setting

## Slide 23: Detector sensitivity to unwanted neutrons (background)
Personnel safety is a necessary but not sufficient function of the shielding (the one investigated by neutronics analysis)
Detector sensitivity to unwanted neutrons is the other function (not investigated by current neutronics analysis)
Experience with the BL4B detector at SNS shows that using B4C to block the detector’s view of concrete surfaces is effective at reducing detector background

## Slide 24: Backdrop must be dark
The backdrop of the sample, viewed by the detector, needs to be as dark as possible. The placement of B4C shielding panels and additional wall-mounted shielding is dictated by this requirement.
QIKR-U
QIKR-L
Every surface painted by the magenta cones is visible to the detector and should be shielded from view using B4C (gold panels)

## Slide 25: Detector assessment
Detector: Timepix3-based scintillator
  Equivalent performance to conventional 3He detector
  Greater (>20×) count-rate capability
  Modular design – replaceable components
  Upgrade path to Timepix4
  → Measuring R down to 10-7 requires low background
Ankner and Khaplanov (STS),
Watkins and Loyd (NScD), and
Long (LANL) data collected at
Lujan Center Asterix reflectometer

## Slide 26: Define beam – reject even more neutrons
Beam-defining slits
Select a satellite beam
Collimate to define resolution and footprint
Reduce intensity >67%
Sample reflectivity
Further reduction of intensity by factor of 10-7-1 depending on λ
→ Shielding in cave must mop up unused thermal and cold neutrons

## Slide 27: There are a lot of neutrons to absorb
Fast neutrons in the monolith and bunker
Thermal and cold neutrons in the instrument cave

## Slide 28: Outline
QIKR, Briefly
Unwanted Neutrons
  Lost in monolith and bunker
  Lost in instrument cave
Summary

## Slide 29: Summary
The curvature of the QIKR guides causes high energy neutrons to depart the guide and enter the bunker→shielding is needed for biological safety and to prevent those neutrons from entering the cave & creating excessive background
The tapering at the end of the QIKR guide causes large neutron loss inside the cave→shielding needs to absorb those neutrons
The tapered guide creates multiple beams exiting the guide→slits select the desired beam and shield (block) the rest
  There is not a reasonable scenario where the slits would be removed before turning on the beam→no useful data can be collected in that state
Additional shielding is needed to minimize unwanted neutrons entering the detector
  An effective solution on BL4B was to have the detector view panels of ZHIP mix (B4C) rather than concrete surfaces→the same will be done on QIKR, though neutronic analysis to confirm its effectiveness has not yet been done

## Slide 30: Thanks
ORNL STS Design Team
John Ankner
Danielle Wilson
Kursat Bekar
Ryan Butz
Joe Griffith
Rudy Thermer
Anton Khaplanov

## Slide 31: Footprint calculator – maximum slits
Resolution
0.02 < δθ/θ < 0.07
implies
0.1 mm < s < 10 mm

## Slide 32: In-bunker components
QIKR-U
Shutter assemblies
B4C / W shutter
Attenuator deck
Beam monitor
Bunker inserts
Monolith insert
Tapered guide
ICS Switches
Bandwidth
choppers
×
QIKR-L

## Slide 33: End-station caves
Elevation
Plan
Isometric
20 m
24 m
Charge 1
Bunker
QIKR-L
QIKR-U

## Slide 34: QIKR Caves & User Hutches – plan view
Bunker Wall
Lower Cave Monorail
Lower Cave Sample Rack
Upper Cave Monorail
Upper Cave Sample Rack
QIKR L
Workstation
Upper Cave End Station
Lower Cave
End Station
IPPS Enclosure
QIKR U
Workstation

## Slide 35: McStas Analysis
McStas can tell the neutrons leave the guide, but can’t tell where they go afterward… the shielding has to mop up those neutrons
MCNP analysis agrees with McStas to within 10% if the same source is used in each
  Therefore MCNP reflects low-energy neutrons correctly
McStas also matches Creo geometry

## Slide 36: Incident Table – holds slits and optics
Preliminary design of an incident table for QIKR
Translating ball screw enables an accurate pivoting in the range ± 5°. Short support for more stability.
A rigid bearing ensures
precise alignment of the slits.
A platform open to the sides, with sufficient distance to the beam line, offers flexible installation space for future installations.
Adjustable machine feet with vibration absorbing pads.
Manually adjustable slit position with safety brake in z-direction.
Adjustability in the xz plane.
Hollow profiles together with energy chain enable protected cable
installation.
The axis of rotation ideally lies on the face and the symmetry plane of the neutron guide.
[courtesy Rudy Thermer]

## Slide 37: Sample Table and Detector Arm unit
Available with different load capacities i.e., various spindle sizes, gear ratios and accuracies.
PI H-850 Hexapod
Payload: 250 kg
FDet
FLift
15°
Preloaded linear ball bushings
ZIMM Screw Jack
With an ideal load-bearing arrangement in relation to the spindle lifting nut.
Increased stiffness
due to a free length
Reduction of 50%.
THK-TU-Series Slide table
providing high load capacity
and accuracy in a single
Structure.
The design employs well-developed standard components (increased reliability, reduced costs).
Nadella GLK (NIRO) Spherical-plain bearing Radial clearance: 0 -10 µm!
Compensatory force transmitting
link.
Hexapods offer six degrees of freedom of movement in the most compact space.
Robust guide system
∅ 60 mm
Yaw swivel unit
Spindle nut
THK Curved linear guides
[courtesy Rudy Thermer]
Charge 1

## Slide 38: Create a layer: Langmuir trough

## Slide 39: Biological Membrane Studies Conducted at Liquids Reflectometer
Lipid membranes form the boundary between cells and the outside world
Membranes (e.g. analogues of lung cells) are spread on a water surface or a solid (e.g. silicon) support
Viral proteins and potential drugs are introduced into the water below the film
Neutrons see the nano-scale rearrangements of the protein relative to the membrane under different conditions
These rearrangements reveal protein function

## Slide 40: Prepare DPPC film and measure NR
Dual barrier KSV NIMA frame assembly
Three trough sizes
Hermetically sealed enclosure
  Sapphire windows for beam transport
  Large viewing area
Active vibration isolation
Trough | Surface Area (cm2) | Subphase Volume (mL) | Compression Ratio
Extra Small | 150 | 18 | 8.7
Small | 98 | 57 | 5.2
Medium | 273 | 176 | 10.8
[Galuska L., Muckley E.S., Cao Z., Ehlenberg D., Qian Z., Zhang S., Rondeau-Gagne S., Phan M.D., Ankner J.F., Ivanov I., Gu X., "SMART transfer method to directly compare the mechanical response of water-supported and free-standing ultrathin polymeric films", Nature Communications 12, 2347 (2021)]

## Slide 41: Langmuir film deposition
KSV 2000 Dual-Barrier Trough
  Large dipping well: 100 x 100 x 35 mm3
  Subphase temperature control
  Total subphase volume: 1.2 L
  Active vibration isolation
KSV NIMA micro BAM
  Fixed angle-of-incidence – 53o
  Large fields of view – 3600μm x 4000μm
  12 μm resolution
  Still images as well as real-time video
μ-BAM image of P3HT-b-PEOT on water subphase
KSV NIMA Surface Potential Meter (SPOT)
  μn = ΔVεε0A
  Determine effective dipole moment
  Determine molecular orientation
  Monitor complex formation between monolayers, subphase species and adsorbates

## Slide 42: Rheology
Anton Paar MCR 501
  Measuring geometries
    Cone-plate (Ti cone w/ single crystal Si base plate)
    Plate-plate (Ti plate w/ single crystal Si base plate)
    Cup and bob ( quartz for RheoSANS and quartz and Al for bench top measurements)
  Temperature control
    Convection (-50 oC to 200 oC)
    Peltier (15 oC to 80 oC)
  RheoSANS
    Bin Wu from BSMD is actively developing a focus area in in situ rheological measurements in SANS
      GP-SANS (HFIR) and EQ-SANS (SNS)

## Slide 43: Multi-Environment Chamber
Chamber base pressure ≤ 10-9 Torr
For surface adsorption studies, absolute gas pressures in the range 0.001 Torr to 1000 Torr.
Radiant resistive heater
  Sample temperatures to 600 oC

## Slide 44: Layer-by-Layer (LbL) growth
Simple
Versatile – vast inventory of polyelectrolytes
Environmentally friendly – aqueous, room temperature
No limitation on substrate shape or size