## Slide 1: Second Generation IRP Design
Presented to the SNS Accelerator Advisory Committee
Jim Janney
Target Systems Engineer
March 9, 2017

## Slide 2: Inner reflector plug overview
The IRP is a 63,000lb plug divided into 3 main segments
  Upper IRP – Carbon steel shielding
  Middle IRP – Water cooled stainless steel shielding
  Lower IRP – Water cooled aluminum vessel filled with steel and beryllium surrounding aluminum ambient and cryogenic moderator vessels
  Water piping and cryogenic transfer lines transverse all three assemblies
Spare
Original

## Slide 3: Lower IRP surrounding structures
Target Module
IRP
Proton Beam

## Slide 4: Moderator locations
Top Upstream Decoupled H2
Top Downstream Coupled H2
Bottom Upstream Decoupled H2O
Bottom Downstream Coupled H2

## Slide 5: Beamline locations
Top Upstream Decoupled H2
Top Downstream Coupled H2
Cd Coated Beamtube
Beamtube without Cd

## Slide 6: Goals for second generation IRP
Simplify Manufacturing
  Spare IRP Manufacturing has taken over 8 years and counting
  Many complicated and difficult to repeat processes
Increase Robustness
  Current IRP is leaking in top downstream water to helium boundary
  Spare IRP is currently in large rebuild effort due to top upstream helium lid weld failure
Increase Lifetime
  Optimize poison/decoupler thicknesses so IRP lifetime is limited by aluminum radiation damage
Maintain Current Physics Design
  Keep same moderator, premoderator, and reflector geometries
  Decision made not to include larger top downstream moderator

## Slide 7: Removal of split plates
Split plates introduce hundreds of aluminum welds
Cooling not needed where aluminum is cooled from at least one side by reflector water flow

## Slide 8: Removal of split plates
Cooling only needed where decoupled (Cd coated) beam ports pass over target port
  Accomplish with deep drilled holes from outside
Remove aluminum between coupled beam ports and target port
Advantages
  Removes lots of aluminum welding
  Stiffens critical structures for aligning target and moderators making it easier to achieve required tolerances
  Removes difficult to analyze and inspect structures

## Slide 9: No cadmium applied to TUS vacuum vessel
Cadmium currently sprayed onto TUS vacuum vessel
  Requires great fit up between vacuum vessel and surrounding structures for cooling
  Risk of damage to vacuum vessel
Instead cadmium will be sprayed onto moderator housing surfaces and helium lid
  Maintains Cd thickness and geometry
  Requires subsequent Cd machining to achieve fit of vacuum vessel
  Less stringent fit up requirement since Cd heat generation does not have to cross fit up gap

## Slide 10: More robust helium lids and lid joints
Helium lids pose significant assembly vulnerabilities as seen in original and spare IRPs
Helium lids can be thickened to increase boundary robustness and make weld joint easier
Weld geometry can be modified for easier access
Considering EB-welding to remove filet welds
Helium Lid Weld Joint

## Slide 11: Transfer lines simplified
Transfer line over constraint is suspected cause in spare IRP helium lid leak
Water and helium layers will be branched off from concentric lines just above intermediate IRP instead of at top of IRP
Changes make transfer lines easier to fabricate and reduce possibility of over constraint
  Slightly higher heat loads to hydrogen systems

## Slide 12: Thickening of gadolinium poison plates
  Current IRP life time limited to about 40 GW-Hrs due to Gadolinium burn up
  IRP life time can be extended to 50 GW-Hrs by thickening poison plates
      Life time cannot be extended further due to radiation damage to aluminum structures closest to target module
      Slight neutron production penalty at beginning of life due to thicker poison plate

## Slide 13: Analysis of design changes is progressing
Structural analysis has verified design changes for pressure and deadweight loadings
Thermal analysis has shown changes are promising, but some modifications to flow distribution baffles have been made to improve cooling
  2nd Iteration CFD analysis on ongoing
Thermal stress analysis will be completed once thermal analysis is completed

## Slide 14: Conclusions
Design for second generation IRP is progressing to simplify manufacturing, make boundaries more robust, and increase lifetime while maintaining the current physics design
  Incorporates lessons learned from Original and Spare IRP fabrication and IRP operations
Structural and thermal analysis shows that design changes are feasible but baffle design has been changed to improve flow distribution and analysis is ongoing